## Taltz SFMEA
Function | Function Status | Sequence of Events Origin | Sequence Number | Sequence of Events | Failure Mode | Software Cause | System Effect | Controls | P1 | Hazardous Situation | P2 | Harms | P-Harm | SEV | Hazard | Risk Level | System Control Verification of Implementation | System Control Verification of Effectiveness
Registration, Login and Profile | Enabled | sFMEA | LTS-1.00 | 1. A new user who has already registered an account within the enrollment workflow and verified their email address is completing the remaining steps during CSP enrollment 2. An API call to the server to complete CSP enrollment fails and returns an error to the calling application 3. The user is confused | Registered user is unable to sign in to use the application | A connection failure occurs between the mobile application and the external server system so the application cannot connect to the server to complete registration | A registered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 5 | 1 | App Configuration | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.01 | 1. User is registered but logged out of the application 2. User opens the application and tries to login 3. User's device is not connected to the internet 4. Application displays an error message that the user is offline |  | The user's device has no network connection so the application cannot connect to the server to complete login | A registered user is unable to use the mobile application. | [RCM-002] The application receives a lost connectivity error message on pages where network connectivity is required only [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 5 | 1 | App Configuration | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.02 | 1. New user tries to submit their data to register an account 2. Application displays an error message that the operation failed 3. User is confused | First time user cannot register or use the app | A connection failure occurs between the mobile application and the external server system so the application cannot connect to the server to complete registration | A first time user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 1 | None | 1 | 1 | User | Low 1-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.03 | 1. New user tries to submit their data to register an account 2. User's device is not connected to the internet 3. Application displays an error message that the operation failed and to connect to the internet 4. User is confused |  | The user's device has no network connection so the application cannot connect to the server to complete registration | A first time user is unable to use the mobile application. | [RCM-002] The application receives a lost connectivity error message on pages where network connectivity is required only | 5 | User dissatisfaction | 1 | None | 1 | 1 | User | Low 1-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.04 | 1. New user opens the application for the first time 2. Application displays an error message that the operation failed since there is no configuration file 3. User is confused |  | App does not check for updated configuration file | A first time user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 1 | None | 1 | 1 | User | Low 1-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.05 | 1. New user tries to open the application for the first time 2. User's device is not connected to the internet 3. Application displays an error message that the operation failed and to connect to the internet 4. User is confused |  | The user's device has no network connection so the application cannot connect to the server to retrieve the configuration file | A first time user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 1 | None | 1 | 1 | User | Low 1-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.06 | 1. User tries to update some data associated with their profile 2. The update fails and application displays an error message 3. User is confused | User is not able to update first name, last name, address, phone number, date of birth, email, or password | Connection failure between client app and external server system | The application displays the data inaccurately to the user. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 4 | 2 | App Configuration | Medium 2-4 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.07 | 1. User tries to update some data associated with their profile 2. Application displays an error message that the operation failed and to connect to the internet 3. User is confused |  | The user's device has no network connection so the application cannot connect to the server to complete registration | The application displays the data inaccurately to the user. | [RCM-002] The application receives a lost connectivity error message on pages where network connectivity is required only [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 4 | 2 | App Configuration | Medium 2-4 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.08 | 1. User tries to update their username or password 2. The update initially succeeds but the email with steps to complete the update is never received 3. User is confused |  | Email not received when resetting password | The application displays the data inaccurately to the user. | [RCM-003] The application will provide an option for the user to resend the email verification email during account registration or when changing the on-file email address. | 5 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 4 | 2 | App Configuration | Medium 2-4 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
App Downloads and App Updates | Enabled | sFMEA | LTS-2.00 | 1. User opens application on a device or operating system version or application version that is no longer supported 2. User sees unsupported version message and cannot use the application and is confused | Existing user is unable to use the app due to using an unsupported device, operating system version or application version | The user's device is unsupported or the operating system version or application version is unsupported | A registered user is unable to use the mobile application. | [RCM-004] The application indicates to the user if they are attempting to use an unsupported device and/or operating system version. [RCM-005] The application indicates to the user if they are attempting to use an unsupported version of the application. | 5 | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 5 | 1 | App Configuration | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
App Configuration | Enabled | sFMEA | LTS-3.00 | 1. User receives the reminder 2. User taps button to manually log an injection 3. User sees automatic injection waiting room screen 4. User administers a dose 5. User does not see dose get automatically logged 6. User administers another dose | An invalid update to the configuration file causes the Automatic Injection Tracking feature for the CAI device to become enabled when it should be disabled. | Valid but incorrect configuration file, including the enablement of a feature which should be disabled, is deployed to the remote server Configuration service and is retrieved by the app | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. [RCM-011] The mobile application will prevent the Automatic Injection Tracking feature from being enabled even if the configuration file received from the remote server has the feature flag enabled. | 1 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 1 | 2 | App Configuration | Low 2-1 | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP RPT-712015 LillyPlus R10 Product Development Traceability Matrix | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 4.05 | 1. User receives the reminder at a different time 2. User has already taken a dose 3. User takes another dose in response to a delayed reminder | An invalid update to the configuration file specific to the dose reminder schedule function results in the user receiving scheduled dose reminders on an incorrect freqeuency | Valid but incorrect configuration file, including an incorrectly configured dose reminder schedule function, is retrieved by the app | Dose reminder is sent on incorrect day | [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. | 3 | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | Dose Reminder | Low 2-2 | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP
|  | sFMEA | LTS-3.01 | 1. The application configuration file is updated on the server with an incorrect language configuration 2. User opens the application 3. Application checks for and retreives updated configuration file 4. Application displays content in a different language 5. User sees different language and delays administering their dose since they cannot determine their scheduled day | An invalid update to the configuration file results in an incorrect language configuration | Valid but incorrect configuration file is retrieved by the app | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. | 3 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | App Configuration | Low 2-2 | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP
|  | sFMEA | LTS-3.02 | 1. The application configuration file is updated on the server with a malformed format 2. User opens the application 3. Application checks for and retreives updated configuration file 4. Application crashes or generally does not perform which prevents the user from being able to use the application | Registered user is unable to load the application content, configuration and features | Invalid configuration is deployed to the cloud Configuration service and retrieved by the app | A registered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. | 3 | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 3 | 1 | App Configuration | Low 1-3 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP
|  | sFMEA | LTS-3.03 | 1. The application configuration file is updated on the server with a malformed format 2. User opens the application 3. Application checks for and retreives updated configuration file 4. Application crashes or generally does not perform which prevents the user from being able to use the application | Unregistered user is unable to load the application content, configuration and features | Invalid configuration is deployed to the cloud Configuration service and retrieved by the app | An unregistered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. | 5 | User dissatisfaction | 1 | None | 1 | 1 | User | Low 1-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP
|  | sFMEA | LTS-3.04 | 1. The application configuration file requires an update 2. The application configuration file is updated on the server 3. User's device is not connected to the internet and does not retrieve the latest configuration file 4. User sees application configuration that is out of date | User sees incorrect feature or content configuration | The user's device has no network connection so the application cannot connect to the server to retrieve configuration file updates. | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-012] After 10 days of offline use, the application displays an alert to the user [RCM-013] After 20 days of offline use, the application displays an alert to the user [RCM-014] After 30 days of offline use, user is prevented from using the application again until they connect to the network [RCM-002] The application receives a lost connectivity error message on pages where network connectivity is required only | 3 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | App Configuration | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-3.05 | 1.The application database encounters a data corruption error and the user's brand program data is lost 2. The application displays generic branding assets for the application instead of product branding assets 3. The user is confused as to why the application looks different |  | Data corruption issue in local application storage causes application to lose brand defined in local data | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 4 | 2 | App Configuration | Medium 2-4 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
Medication Dose Reminders | Enabled | aFMEA | 4.04 | 1. User receives the reminder at a different time 2. User decides to not take the dose 3. User misses the dose | Medication dose reminder is not delivered | Operating system fails to deliver scheduled reminder notification | Dose reminder is not sent | [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-017] If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. [RCM-018] If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-019] If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. | 2 | No medication reminder is sent | 4 | Less than full therapeutic benefit (Underdose) | 1 | 2 | Dose Reminder | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | Application fails to schedule reminder notification | Dose reminder is not sent | [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-017] If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. [RCM-018] If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-019] If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. | 2 | No medication reminder is sent | 4 | Less than full therapeutic benefit (Underdose) | 1 | 2 | Dose Reminder | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 4.01 | 1. User receives the reminder 2. User ignores or does not acknowledge (click) the reminder 3. User performs the injection 4. User still sees the reminder open 5. User decides to take another dose | Too many medication dose reminders are sent | Application schedules too many reminder notifications | Incorrect medication reminders are sent | [RCM-017] If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. [RCM-018] If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-019] If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-020] The application will prevent the user from logging more than 2 injections on any given day. [RCM-021] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user before logging the injection. [RCM-022] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user while logging the injection. [RCM-023] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user after logging the injection. | 3 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | App Configuration | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 4.06 | 1. User receives the reminder at a different time 2. User was not expecting it and gets frustrated | Medication dose reminder is sent at the incorrect time on the correct day | Application schedules reminder notification incorrectly | Dose reminder is sent at incorrect time | None | 5 | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 4 | 2 | Dose Reminder | Medium 2-4 | N/A | N/A
[6 empty] | Operating delivers reminder notification incorrectly | Dose reminder is sent at incorrect time | None | 5 | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 4 | 2 | Dose Reminder | Medium 2-4 | N/A | N/A
|  | aFMEA | 4.05 | 1. User receives the reminder at a different time 2. User has already taken a dose 3. User takes another dose in response to a delayed reminder | Medication dose reminder is sent on the incorrect day | Application schedules reminder notification incorrectly | Dose reminder is sent on incorrect day | [RCM-017] If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. [RCM-018] If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-019] If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-020] The application will prevent the user from logging more than 2 injections on any given day. [RCM-021] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user before logging the injection. [RCM-022] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user while logging the injection. [RCM-023] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user after logging the injection. [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-035] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. In-application prompts which also display the scheduled dose date will display the correct scheduled dose date regardless of a time zone change or daylight savings change. [RCM-036] The medication prescribing information is linked to from within the application. | 3 | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | Dose Reminder | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | PLN-66872 LillyPlus SaMD R10.0.1 Emergency Change Plan (MMA-279) RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-4.01 | 1. User has created their treatment plan 2. User's device has changed time zone due to a Daylight Savings Time change 3. User sees an incorrect scheduled dose day on the Primary Action Card within the application | Medication in-app prompt on the Primary Action Card is displaying the incorrect dosing day | Application calculates the dosing day on the Primary Action Card incorrectly | Application displays an incorrect scheduled dosing day | [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-020] The application will prevent the user from logging more than 2 injections on any given day. [RCM-021] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user before logging the injection. [RCM-022] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user while logging the injection. [RCM-023] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user after logging the injection. [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-035] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. In-application prompts which also display the scheduled dose date will display the correct scheduled dose date regardless of a time zone change or daylight savings change. [RCM-036] The medication prescribing information is linked to from within the application. | 3 | Incorrect scheduled dosing day is identified | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | Dose Logging | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | PLN-66872 LillyPlus SaMD R10.0.1 Emergency Change Plan (MMA-279) RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-4.00 | 1. User begins treatment plan setup 2. Treatment plan setup fails 3. User is dissatisfied | User cannot set up treatment plan | A connection failure occurs between the mobile application and the external server system so the application cannot connect to the server to complete registration | A registered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 5 | 1 | App Configuration | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | The data inputs fail to save to the database | A registered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 5 | 1 | App Configuration | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
Symptom Tracking | Enabled | aFMEA | 6.04 | 1. User reviews logbook 2. User decides to use the data for their treatment 3. User does not talk to doctor and changes dosing regimen or medications | Misrepresentation of data in graphs | LillyCloud framework causes data sent to the calling application to display inaccurately | App presents incorrect logbook data | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-024] Symptoms are represented on a bar graph by displaying the raw values the user recorded for each symptom for each day of the week. There are no calculations involved in deriving the data that displays on the graph. | 4 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 3 | 2 | App Configuration | Low 2-3 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 6.01 | 1. User reviews logbook with inaccurate data 2. User decides to use the data for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Data loss | Data not yet synchronized to the server is lost when a user switches to a new primary device | App presents incorrect logbook data | [RCM-025] If the user attempts to use the application on a device that is not the primary device, the user will verify intent to register the device as primary via the primary device check screen before dis-associating the prior device. This alert message informs the user that any local data that hasn't been synced to the server will be lost [RCM-026] When the application connects to the Digital Health cloud, any symptom tracking entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-027] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical symptom tracking data of the user will be downloaded from the cloud to the device. [RCM-028] When a user scrolls past their 6 months of downloaded symptom tracking data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of symptom data. | 2 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 1 | 2 | App Configuration | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-5.00 | 1. User record symptoms with a photo 2. Photo appears in user's device camera roll 3. User is dissatisfied | Photos saved to device camera roll / photo gallery instead of only the application | Change in operating system functionality causes tracked photos to be recorded in the operating system camera roll rather than the application | User's tracked symptom photos are stored in the device camera roll | None | 5 | User dissatisfaction | 1 | None | 1 | 1 | User | Low 1-1 | N/A | N/A
Content Resources | Enabled | sFMEA | LTS-6.00 | 1. User reviews content resource with incorrect treatment content 2. User devices to use this information for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Unaproved content resource in app | Configuration error | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. | 3 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | App Configuration | Low 2-2 | PRD-40196 — Digital Health - Content and Configuration Management SOP | PRD-40196 — Digital Health - Content and Configuration Management SOP
[6 empty] | Existing Lilly owned content is modified by another Lilly team prior to or without approval | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-029] Changes to any application content must go through the appropriate content review process (i.e. PCA / SCRAT / MLRO). The application product team becomes aware of proposed changes to content approved via these processes. | 3 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | App Configuration | Low 2-2 | PDS-SOP-00344 Software Content Review and Approval Team (SCRAT) USCIQ-SOP-221-001: General Requirements for All Tactics | PDS-SOP-00344 Software Content Review and Approval Team (SCRAT) USCIQ-SOP-221-001: General Requirements for All Tactics
Automatic Injection Tracking | Function Disabled for R11 Commercial Release | aFMEA | 6.01 | 1. User reviews logbook with inaccurate data 2. User decides to use the data for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Injection received automatically via Bluetooth is unable to be logged to the logbook | Decryption of injection data fails due to error in connection with Lilly cloud while retrieving decryption key for the device, resulting in no injection being logged to the logbook since the data cannot be authenticated as coming from a Lilly device | App does not log successful injection event | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-031] When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-030] The application provides the user the option to manually log their injection if the Connected Taltz Delivery System identifies an injection error. | 2 | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 1 | 2 | Dose Logging | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | Decryption of injection data fails due to an incorrectly formatted request from the mobile application, resulting in no injection being logged to the logbook since the data cannot be authenticated as coming from a Lilly device | App does not log successful injection event | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-030] The application provide the user the option to manually log their injection if the Connected Taltz Delivery System identifies an injection error. | 2 | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 1 | 2 | Dose Logging | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | Decryption of injection data fails due to LilyAutoInjector SDK not working properly, resulting in no injection being logged to the logbook since the data cannot be authenticated as coming from a Lilly device | App does not log successful injection event | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-030] The application provide the user the option to manually log their injection if the Connected Taltz Delivery System identifies an injection error. | 2 | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 1 | 2 | Dose Logging | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | Decryption of injection data fails due to the device having no network connection, resulting in no injection being logged to the logbook since the data cannot be authenticated as coming from a Lilly device | App does not log successful injection event | [RCM-031] When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-032] If the application receives injection data from a nearby compatible version of the Connected Taltz Delivery System but is not connected to the internet, the application will display an offline error message. [RCM-030] The application provide the user the option to manually log their injection if the Connected Taltz Delivery System identifies an injection error. | 2 | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 1 | 2 | Dose Logging | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | Corruption of data in local application storage causes loss of received injection data | App does not log successful injection event | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-030] The application provide the user the option to manually log their injection if the Connected Taltz Delivery System identifies an injection error. | 2 | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 1 | 2 | Dose Logging | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
Manual Injection Tracking | Enabled | aFMEA | 6.01 | 1. User reviews logbook with inaccurate data 2. User decides to use the data for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Inaccurate information is displayed to the user | Connection failure between client app and external system | App presents incorrect logbook data | [RCM-031] When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-033] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical medication administration logging data of the user will be downloaded from the cloud to the device. [RCM-034] When a user scrolls past their 6 months of downloaded injection logging data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of medication administration data. | 3 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | App Configuration | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | LillyCloud framework causes data sent to the calling application to display inaccurately | App presents incorrect logbook data | [RCM-031] When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-033] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical medication administration logging data of the user will be downloaded from the cloud to the device. [RCM-034] When a user scrolls past their 6 months of downloaded injection logging data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of medication administration data. | 3 | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | 2 | App Configuration | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 6.01 | 1. User reviews logbook with inaccurate data 2. User decides to use the data for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Data loss | Data not yet synchronized to the server is lost when a user switches to a new primary device | App does not log successful injection event | [RCM-025] If the user attempts to use the application on a device that is not the primary device, the user will verify intent to register the device as primary via the primary device check screen before dis-associating the prior device. This alert message informs the user that any local data that hasn't been synced to the server will be lost [RCM-031] When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-033] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical medication administration logging data of the user will be downloaded from the cloud to the device. [RCM-034] When a user scrolls past their 6 months of downloaded injection logging data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of medication administration data. | 2 | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 1 | 2 | Dose Logging | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report

## Omvoh SFMEA
Function | Function Status | Sequence of Events Origin | Sequence Number | Sequence of Events | Failure Mode | Software Cause | System Effect | Controls | P1 | Hazardous Situation | P2 | Harms | P-Harm | SEV | Hazard | Risk Level | System Control Verification of Implementation | System Control Verification of Effectiveness
Registration, Login and Profile | Enabled | sFMEA | LTS-1.00 | 1. A new user who has already registered an account within the enrollment workflow and verified their email address is completing the remaining steps during CSP enrollment 2. An API call to the server to complete CSP enrollment fails and returns an error to the calling application 3. The user is confused | Registered user is unable to sign in to use the application | A connection failure occurs between the mobile application and the external server system so the application cannot connect to the server to complete registration | A registered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.01 | 1. User is registered but logged out of the application 2. User opens the application and tries to login 3. User's device is not connected to the internet 4. Application displays an error message that the user is offline |  | The user's device has no network connection so the application cannot connect to the server to complete login | A registered user is unable to use the mobile application. | [RCM-002] The application receives a lost connectivity error message on pages where network connectivity is required only [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.02 | 1. New user tries to submit their data to register an account 2. Application displays an error message that the operation failed 3. User is confused | First time user cannot register or use the app | A connection failure occurs between the mobile application and the external server system so the application cannot connect to the server to complete registration | A first time user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.03 | 1. New user tries to submit their data to register an account 2. User's device is not connected to the internet 3. Application displays an error message that the operation failed and to connect to the internet 4. User is confused |  | The user's device has no network connection so the application cannot connect to the server to complete registration | A first time user is unable to use the mobile application. | [RCM-002] The application receives a lost connectivity error message on pages where network connectivity is required only | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.04 | 1. New user opens the application for the first time 2. Application displays an error message that the operation failed since there is no configuration file 3. User is confused |  | App does not check for updated configuration file | A first time user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.05 | 1. New user tries to open the application for the first time 2. User's device is not connected to the internet 3. Application displays an error message that the operation failed and to connect to the internet 4. User is confused |  | The user's device has no network connection so the application cannot connect to the server to retrieve the configuration file | A first time user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.06 | 1. User tries to update some data associated with their profile 2. The update fails and application displays an error message 3. User is confused | User is not able to update first name, last name, address, phone number, date of birth, email, or password | Connection failure between client app and external server system | The application displays the data inaccurately to the user. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.07 | 1. User tries to update some data associated with their profile 2. Application displays an error message that the operation failed and to connect to the internet 3. User is confused |  | The user's device has no network connection so the application cannot connect to the server to complete registration | The application displays the data inaccurately to the user. | [RCM-002] The application receives a lost connectivity error message on pages where network connectivity is required only [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-1.08 | 1. User tries to update their username or password 2. The update initially succeeds but the email with steps to complete the update is never received 3. User is confused |  | Email not received when resetting password | The application displays the data inaccurately to the user. | [RCM-003] The application will provide an option for the user to resend the email verification email during account registration or when changing the on-file email address. | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
App Downloads and App Updates | Enabled | sFMEA | LTS-2.00 | 1. User opens application on a device or operating system version or application version that is no longer supported 2. User sees unsupported version message and cannot use the application and is confused | Existing user is unable to use the app due to using an unsupported device, operating system version or application version | The user's device is unsupported or the operating system version or application version is unsupported | A registered user is unable to use the mobile application. | [RCM-004] The application indicates to the user if they are attempting to use an unsupported device and/or operating system version. [RCM-005] The application indicates to the user if they are attempting to use an unsupported version of the application. | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
App Configuration | Enabled | sFMEA | LTS-3.00 | 1. User receives the reminder 2. User taps button to manually log an injection 3. User sees automatic injection waiting room screen 4. User administers a dose 5. User does not see dose get automatically logged 6. User administers another dose | An invalid update to the configuration file causes the Automatic Injection Tracking feature for the CAI device to become enabled when it should be disabled. | Valid but incorrect configuration file, including the enablement of a feature which should be disabled, is deployed to the remote server Configuration service and is retrieved by the app | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. [RCM-011] The mobile application will prevent the Automatic Injection Tracking feature from being enabled even if the configuration file received from the remote server has the feature flag enabled. | 1 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 1 | 2 | Dose | Low 2-1 | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP RPT-712015 LillyPlus R10 Product Development Traceability Matrix | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 4.05 | 1. User receives the reminder at a different time 2. User has already taken a dose 3. User takes another dose in response to a delayed reminder | An invalid update to the configuration file specific to the dose reminder schedule function results in the user receiving scheduled dose reminders on an incorrect freqeuency | Valid but incorrect configuration file, including an incorrectly configured dose reminder schedule function, is retrieved by the app | Dose reminder is sent on incorrect day | [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP
|  | sFMEA | LTS-3.01 | 1. The application configuration file is updated on the server with an incorrect language configuration 2. User opens the application 3. Application checks for and retreives updated configuration file 4. Application displays content in a different language 5. User sees different language and delays administering their dose since they cannot determine their scheduled day | An invalid update to the configuration file results in an incorrect language configuration | Valid but incorrect configuration file is retrieved by the app | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. | 3 | Delay in therapy (device not used) | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact | 3 | 1 | Dose | Low 1-3 | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP | PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP
|  | sFMEA | LTS-3.02 | 1. The application configuration file is updated on the server with a malformed format 2. User opens the application 3. Application checks for and retreives updated configuration file 4. Application crashes or generally does not perform which prevents the user from being able to use the application | Registered user is unable to load the application content, configuration and features | Invalid configuration is deployed to the cloud Configuration service and retrieved by the app | A registered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. | 3 | Delay in therapy (device not used) | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact | 3 | 1 | Dose | Low 1-3 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP
|  | sFMEA | LTS-3.03 | 1. The application configuration file is updated on the server with a malformed format 2. User opens the application 3. Application checks for and retreives updated configuration file 4. Application crashes or generally does not perform which prevents the user from being able to use the application | Unregistered user is unable to load the application content, configuration and features | Invalid configuration is deployed to the cloud Configuration service and retrieved by the app | An unregistered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-006] The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. [RCM-008] The SOP defines the process for how updates to the configuration file are deployed. [RCM-009] The SOP defines the security plan for the platform on which the configuration files reside. [RCM-010] The SOP defines the security procedures for the platform on which the configuration files reside. | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report PLN-52090 LillyPlus Maintenance Plan PRD-40196 — Digital Health - Content and Configuration Management SOP ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP ITLIB-CSV-015452 Digital Health AWS Platform Security Plan PRD-11289 Digital Health AWS Security Administration and Support SOP
|  | sFMEA | LTS-3.04 | 1. The application configuration file requires an update 2. The application configuration file is updated on the server 3. User's device is not connected to the internet and does not retrieve the latest configuration file 4. User sees application configuration that is out of date | User sees incorrect feature or content configuration | The user's device has no network connection so the application cannot connect to the server to retrieve configuration file updates. | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-012] After 10 days of offline use, the application displays an alert to the user [RCM-013] After 20 days of offline use, the application displays an alert to the user [RCM-014] After 30 days of offline use, user is prevented from using the application again until they connect to the network [RCM-002] The application receives a lost connectivity error message on pages where network connectivity is required only | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-3.05 | 1.The application database encounters a data corruption error and the user's brand program data is lost 2. The application displays generic branding assets for the application instead of product branding assets 3. The user is confused as to why the application looks different |  | Data corruption issue in local application storage causes application to lose brand defined in local data | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
Medication Dose Reminders | Enabled | aFMEA | 4.04 | 1. User receives the reminder at a different time 2. User decides to not take the dose 3. User misses the dose | Medication dose reminder is not delivered | Operating system fails to deliver scheduled reminder notification | Dose reminder is not sent | [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-017] If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. [RCM-018] If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-019] If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. | 2 | Underdose | 4 | Minor: Less than full therapeutic benefit | 1 | 2 | Dose | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | Application fails to schedule reminder notification | Dose reminder is not sent | [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-017] If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. [RCM-018] If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-019] If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. | 2 | Underdose | 4 | Minor: Less than full therapeutic benefit | 1 | 2 | Dose | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 4.01 | 1. User receives the reminder 2. User ignores or does not acknowledge (click) the reminder 3. User performs the injection 4. User still sees the reminder open 5. User decides to take another dose | Too many medication dose reminders are sent | Application schedules too many reminder notifications | Incorrect medication reminders are sent | [RCM-017] If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. [RCM-018] If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-019] If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-020] The application will prevent the user from logging more than 2 injections on any given day. [RCM-021] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user before logging the injection. [RCM-022] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user while logging the injection. [RCM-023] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user after logging the injection. | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 4.06 | 1. User receives the reminder at a different time 2. User was not expecting it and gets frustrated | Medication dose reminder is sent at the incorrect time on the correct day | Application schedules reminder notification incorrectly | Dose reminder is sent at incorrect time | None | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | N/A | N/A
[6 empty] | Operating delivers reminder notification incorrectly | Dose reminder is sent at incorrect time | None | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | N/A | N/A
|  | aFMEA | 4.05 | 1. User receives the reminder at a different time 2. User has already taken a dose 3. User takes another dose in response to a delayed reminder | Medication dose reminder is sent on the incorrect day | Application schedules reminder notification incorrectly | Dose reminder is sent on incorrect day | [RCM-017] If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. [RCM-018] If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-019] If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-020] The application will prevent the user from logging more than 2 injections on any given day. [RCM-021] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user before logging the injection. [RCM-022] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user while logging the injection. [RCM-023] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user after logging the injection. [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-035] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. In-application prompts which also display the scheduled dose date will display the correct scheduled dose date regardless of a time zone change or daylight savings change. [RCM-036] The medication prescribing information is linked to from within the application. | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | PLN-66872 LillyPlus SaMD R10.0.1 Emergency Change Plan (MMA-279) RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-4.01 | 1. User has created their treatment plan 2. User's device has changed time zone due to a Daylight Savings Time change 3. User sees an incorrect scheduled dose day on the Primary Action Card within the application | Medication in-app prompt on the Primary Action Card is displaying the incorrect dosing day | Application calculates the dosing day on the Primary Action Card incorrectly | Application displays an incorrect scheduled dosing day | [RCM-016] The user's dosing schedule is available within the Plan section of the app [RCM-020] The application will prevent the user from logging more than 2 injections on any given day. [RCM-021] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user before logging the injection. [RCM-022] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user while logging the injection. [RCM-023] If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user after logging the injection. [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-035] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. In-application prompts which also display the scheduled dose date will display the correct scheduled dose date regardless of a time zone change or daylight savings change. [RCM-036] The medication prescribing information is linked to from within the application. | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | PLN-66872 LillyPlus SaMD R10.0.1 Emergency Change Plan (MMA-279) RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-4.00 | 1. User begins treatment plan setup 2. Treatment plan setup fails 3. User is dissatisfied | User cannot set up treatment plan | A connection failure occurs between the mobile application and the external server system so the application cannot connect to the server to complete registration | A registered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | The data inputs fail to save to the database | A registered user is unable to use the mobile application. | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
Symptom Tracking | Enabled | aFMEA | 6.04 | 1. User reviews logbook 2. User decides to use the data for their treatment 3. User does not talk to doctor and changes dosing regimen or medications | Misrepresentation of data in graphs | LillyCloud framework causes data sent to the calling application to display inaccurately | App presents incorrect logbook data | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support [RCM-024] Symptoms are represented on a bar graph by displaying the raw values the user recorded for each symptom for each day of the week. There are no calculations involved in deriving the data that displays on the graph. | 4 | Underdose | 4 | Minor: Less than full therapeutic benefit | 3 | 2 | Dose | Low 2-3 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 6.01 | 1. User reviews logbook with inaccurate data 2. User decides to use the data for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Data loss | Data not yet synchronized to the server is lost when a user switches to a new primary device | App presents incorrect logbook data | [RCM-025] If the user attempts to use the application on a device that is not the primary device, the user will verify intent to register the device as primary via the primary device check screen before dis-associating the prior device. This alert message informs the user that any local data that hasn't been synced to the server will be lost [RCM-026] When the application connects to the Digital Health cloud, any symptom tracking entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-027] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical symptom tracking data of the user will be downloaded from the cloud to the device. [RCM-028] When a user scrolls past their 6 months of downloaded symptom tracking data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of symptom data. | 2 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 1 | 2 | Dose | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | sFMEA | LTS-5.00 | 1. User record symptoms with a photo 2. Photo appears in user's device camera roll 3. User is dissatisfied | Photos saved to device camera roll / photo gallery instead of only the application | Change in operating system functionality causes tracked photos to be recorded in the operating system camera roll rather than the application | User's tracked symptom photos are stored in the device camera roll | None | 5 | User dissatisfaction | 5 | User dissatisfied | 5 | 1 | User Experience | Low 1-5 | N/A | N/A
Content Resources | Enabled | sFMEA | LTS-6.00 | 1. User reviews content resource with incorrect treatment content 2. User devices to use this information for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Unaproved content resource in app | Configuration error | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-007] The SOP defines the process for the proper management of the application content and configuration file. | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | PRD-40196 — Digital Health - Content and Configuration Management SOP | PRD-40196 — Digital Health - Content and Configuration Management SOP
[6 empty] | Existing Lilly owned content is modified by another Lilly team prior to or without approval | Incorrect configuration is used causing incorrect features or content to be displayed | [RCM-029] Changes to any application content must go through the appropriate content review process (i.e. PCA / SCRAT / MLRO). The application product team becomes aware of proposed changes to content approved via these processes. | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | PDS-SOP-00344 Software Content Review and Approval Team (SCRAT) USCIQ-SOP-221-001: General Requirements for All Tactics | PDS-SOP-00344 Software Content Review and Approval Team (SCRAT) USCIQ-SOP-221-001: General Requirements for All Tactics
Manual Injection Tracking | Enabled | aFMEA | 6.01 | 1. User reviews logbook with inaccurate data 2. User decides to use the data for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Inaccurate information is displayed to the user | Connection failure between client app and external system | App presents incorrect logbook data | [RCM-031] When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-033] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical medication administration logging data of the user will be downloaded from the cloud to the device. [RCM-034] When a user scrolls past their 6 months of downloaded injection logging data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of medication administration data. | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
[6 empty] | LillyCloud framework causes data sent to the calling application to display inaccurately | App presents incorrect logbook data | [RCM-031] When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-033] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical medication administration logging data of the user will be downloaded from the cloud to the device. [RCM-034] When a user scrolls past their 6 months of downloaded injection logging data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of medication administration data. | 3 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | 2 | Dose | Low 2-2 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
|  | aFMEA | 6.01 | 1. User reviews logbook with inaccurate data 2. User decides to use the data for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Data loss | Data not yet synchronized to the server is lost when a user switches to a new primary device | App does not log successful injection event | [RCM-025] If the user attempts to use the application on a device that is not the primary device, the user will verify intent to register the device as primary via the primary device check screen before dis-associating the prior device. This alert message informs the user that any local data that hasn't been synced to the server will be lost [RCM-031] When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-033] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical medication administration logging data of the user will be downloaded from the cloud to the device. [RCM-034] When a user scrolls past their 6 months of downloaded injection logging data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of medication administration data. | 2 | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 1 | 2 | Dose | Low 2-1 | RPT-712015 LillyPlus R10 Product Development Traceability Matrix | RPT-816867 LillyPlus R10 Design Verification and Validation Summary Report
Accident Tracking | Enabled | aFMEA | 6.04 | 1. User reviews logbook 2. User decides to use the data for their treatment 3. User does not talk to doctor and changes dosing regimen or medications | Misrepresentation of data in graphs | LillyCloud framework causes data sent to the calling application to display inaccurately | App presents incorrect logbook data | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support
|  | aFMEA | 6.01 | 1. User reviews logbook with inaccurate data 2. User decides to use the data for their regimen 3. User uses the drug product inaccurately 4. User takes another injection | Data loss | Data not yet synchronized to the server is lost when a user switches to a new primary device | App presents incorrect logbook data | [RCM-025] If the user attempts to use the application on a device that is not the primary device, the user will verify intent to register the device as primary via the primary device check screen before dis-associating the prior device. This alert message informs the user that any local data that hasn't been synced to the server will be lost [RCM-037] When the application connects to the Digital Health cloud, any accident tracking entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. [RCM-038] When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical accident tracking data of the user will be downloaded from the cloud to the device. [RCM-039] When a user scrolls past their 6 months of downloaded accident tracking data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of accident data.
Infusion Tracking and Infusion Appointment Reminders | Enabled | aFMEA | 6.04 | 1. User reviews logbook 2. User decides to use the data for their treatment 3. User does not talk to doctor and changes dosing regimen or medications | Misrepresentation of data in graphs | LillyCloud framework causes data sent to the calling application to display inaccurately | App presents incorrect logbook data | [RCM-001] Error message alerting the user that the operation failed and to try again or contact support
|  | aFMEA | 6.04 | 1. User reviews logbook 2. User decides to use the data for their treatment 3. User does not talk to doctor and changes dosing regimen or medications | Inaccurate information is displayed to the user | Connection failure between client app and external system | App presents incorrect logbook data | [RCM-037] When the application connects to the Digital Health cloud, any accident tracking entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced.
[6 empty] | LillyCloud framework causes data sent to the calling application to display inaccurately | App presents incorrect logbook data | [RCM-037] When the application connects to the Digital Health cloud, any accident tracking entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced.
|  | aFMEA | 6.04 | 1. User reviews logbook 2. User decides to use the data for their treatment 3. User does not talk to doctor and changes dosing regimen or medications | Data loss | Data not yet synchronized to the server is lost when a user switches to a new primary device | App does not log successful injection event | [RCM-025] If the user attempts to use the application on a device that is not the primary device, the user will verify intent to register the device as primary via the primary device check screen before dis-associating the prior device. This alert message informs the user that any local data that hasn't been synced to the server will be lost [RCM-031] When the application connects to the Digital Health cloud, any logged injection entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced.
|  | aFMEA | 4.04 | 1. User receives the reminder at a different time 2. User decides to not take the dose 3. User misses the dose | Infusion appointment reminder is not delivered | Operating system fails to deliver scheduled reminder notification | User-scheduled appointment reminder is not sent | [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-016] The user's dosing schedule is available within the Plan section of the app
[6 empty] | Application fails to schedule reminder notification | User-scheduled appointment reminder is not sent | [RCM-015] Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. [RCM-016] The user's dosing schedule is available within the Plan section of the app

## RCMs
RCM ID | RCM Text | Control Requirement
RCM-001 | Error message alerting the user that the operation failed and to try again or contact support | LYP-STRY-446: Registration Error Messages
RCM-002 | The application receives a lost connectivity error message on pages where network connectivity is required only | LYP-STRY-73: Lost Connectivity
RCM-003 | The application will provide an option for the user to resend the email verification email during account registration or when changing the on-file email address. | LYP-STRY-558: Resend Email Verification Email
RCM-004 | The application indicates to the user if they are attempting to use an unsupported device and/or operating system version. | LYP-STRY-143: Minimum required OS or Device
RCM-005 | The application indicates to the user if they are attempting to use an unsupported version of the application. | LYP-STRY-142: The application shall support forced upgrades
RCM-006 | The LillyPlus Maintenance Plan includes a design transfer checklist deliverable which is completed during and after the deployment activities for a release. The design transfer checklist includes a verification step to ensure the configuration file is correctly deployed. | PLN-52090 LillyPlus Maintenance Plan
RCM-007 | The SOP defines the process for the proper management of the application content and configuration file. | PRD-40196 Digital Health - Content and Configuration Management SOP
RCM-008 | The SOP defines the process for how updates to the configuration file are deployed. | ITLIB-SOP-000694 Digital Health Config Portal System Admin SOP
RCM-009 | The SOP defines the security plan for the platform on which the configuration files reside. | ITLIB-CSV-015452 Digital Health AWS Platform Security Plan
RCM-010 | The SOP defines the security procedures for the platform on which the configuration files reside. | PRD-11289 — Digital Health AWS Security Administration and Support SOP
RCM-011 | The mobile application will prevent the Automatic Injection Tracking feature from being enabled even if the configuration file received from the remote server has the feature flag enabled. | LYP-STRY-512: Automatic Injection Logging Configurability
RCM-012 | After 10 days of offline use, the application displays an alert to the user | LYP-STRY-544: Offline for 10 Days
RCM-013 | After 20 days of offline use, the application displays an alert to the user | LYP-STRY-545: Offline for 20 Days
RCM-014 | After 30 days of offline use, user is prevented from using the application again until they connect to the network | LYP-STRY-546: Offline for 30 Days
RCM-015 | Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. | LYP-STRY-554: Time Zone Change for Dose Reminders
RCM-016 | The user's dosing schedule is available within the Plan section of the app | LYP-STRY-490: Medication Treatment Plan
RCM-017 | If there is a scheduled dose for the current day, the application will display a primary action card on the Home screen for the user to log a dose. | LYP-STRY-502: Primary Action Card 4 - Dose Day (Not First Dose)
RCM-018 | If the user had a scheduled dose yesterday but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. | LYP-STRY-503: Primary Action Card 5 - Dose Day + 1
RCM-019 | If the user had a scheduled dose 2 or 3 days prior but did not log it, the application will display a primary action card on the Home screen for the user to log their dose. | LYP-STRY-504: Primary Action Card 6 - Dose Day + 2 or 3
RCM-020 | The application will prevent the user from logging more than 2 injections on any given day. | LYP-STRY-536: Maximum Taltz Injections Logged Per Day
RCM-021 | If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user before logging the injection. | LYP-STRY-473: Taltz Manual Schedule Adherence Messaging 1 Pre-Logging
RCM-022 | If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user while logging the injection. | LYP-STRY-474: Taltz Manual Schedule Adherence Messaging 2 During Logging
RCM-023 | If the user manually logs an injection on a day that is not their scheduled dose day or manually logs an extra injection when they have already logged the appropriate number of injection(s) (per the treatment plan) on a scheduled dose day, the application provides schedule adherence messaging to the user after logging the injection. | LYP-STRY-475: Taltz Manual Schedule Adherence Messaging 3 After Logging
RCM-024 | Symptoms are represented on a bar graph by displaying the raw values the user recorded for each symptom for each day of the week. There are no calculations involved in deriving the data that displays on the graph. | LYP-STRY-182: Progress Symptom Tracking
RCM-025 | If the user attempts to use the application on a device that is not the primary device, the user will verify intent to register the device as primary via the primary device check screen before dis-associating the prior device. This alert message informs the user that any local data that hasn't been synced to the server will be lost | LYP-STRY-395: New Primary Device
RCM-026 | When the application connects to the Digital Health cloud, any symptom tracking entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. | LYP-STRY-421: Sync Symptom Data to Cloud
RCM-027 | When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical symptom tracking data of the user will be downloaded from the cloud to the device. | LYP-STRY-525: Symptom Tracking Historical Data
RCM-028 | When a user scrolls past their 6 months of downloaded symptom tracking data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of symptom data. | LYP-STRY-526: Symptom Tracking Additional Historical Data
RCM-029 | Changes to any application content must go through the appropriate content review process (i.e. PCA / SCRAT / MLRO). The application product team becomes aware of proposed changes to content approved via these processes. | PDS-SOP-00344: SOFTWARE CONTENT REVIEW AND APPROVAL TEAM (SCRAT), USCIQ-SOP-221-001: General Requirements for All Tactics
RCM-030 | The application provide the user the option to manually log their injection if the Connected Taltz Delivery System identifies an injection error. | LYP-STRY-540: Taltz Automatic Injection Logging Error
RCM-031 | When the application connects to the Digital Health cloud, any logged medication administration entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced. | LYP-STRY-557: Injection Data Sync to Cloud
RCM-032 | If the application receives injection data from a nearby compatible version of the Connected Taltz Delivery System but is not connected to the internet, the application will display an offline error message. | LYP-STRY-521: Automatic Injection Logging - No Internet Scenario
RCM-033 | When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical medication administration logging data of the user will be downloaded from the cloud to the device. | LYP-STRY-527: Injection Logging Historical Data
RCM-034 | When a user scrolls past their 6 months of downloaded injection logging data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of medication administration data. | LYP-STRY-528: Injection Logging Additional Historical Data
RCM-035 | Scheduled dose reminder notifications will always be delivered at the time requested by the user in the local time zone, including when a device time zone change has been detected. The scheduled notification is not modified by the application as a result of a time zone change. In-application prompts which also display the scheduled dose date will display the correct scheduled dose date regardless of a time zone change or daylight savings change. | LYP-STRY-561: Time Zone Change for Dose Reminders and In-App Prompts
RCM-036 | The medication prescribing information is linked to from within the application. | LYP-STRY-490: Medication Treatment Plan
RCM-037 | When the application connects to the Digital Health cloud, any accident tracking entries, changes, or deletions on the mobile device will be uploaded if they have not already been uploaded/synced.
RCM-038 | When the user logs in on a new device and sets the device as the primary device, the profile data of the user and up to 6 months of the most recent historical accident tracking data of the user will be downloaded from the cloud to the device.
RCM-039 | When a user scrolls past their 6 months of downloaded accident tracking data after a new primary device sync, the application will request additional historical data from the Lilly cloud up to the earliest logged date of accident data.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: same as above

## PHA_ORIG
Hazard | Hazardous Situation | P2 | Harm | Severity | Regulatory Review | Comments
User Experience | User dissatisfaction | 5 | User dissatisfied | 1 | No | None
No Hazard | N/A - other situation with no associated hazard | N/A | N/A – no clinical impact | N/A | No | None
Dose | Repeated missed dose or partial dose for up to 3 months | 4 | Minor: Less than full therapeutic benefit, possible requiring medical intervention | 3 | Yes | P2 reflects variability in user response / sensitivity to repeated underdose and variability in the % underdose
Dose | Delivery of drug with less than full potency repeated for up to 3 months (equivalent to underdose) | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention | 3 | Yes | P2 reflects variability in user response / sensitivity to repeated underdose and variability in the % underdose
Dose | Underdose | 4 | Minor: Less than full therapeutic benefit | 2 | Yes | P2 estimate reflects that most minor underdoses are not symptomatic or provide ample opportunity for users to correct before significant impact
Dose | No dose delivered | 4 | Minor: Less than full therapeutic benefit | 2 | Yes | P2 estimate reflects that most minor underdoses are not symptomatic or provide ample opportunity for users to correct before significant impact
Dose | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | Yes | P2 estimate reflects that most minor underdoses are not symptomatic or provide ample opportunity for users to correct before significant impact
Dose | Delivery of drug with less than full potency (equivalent to underdose) | 4 | Minor: Less than full therapeutic benefit | 2 | Yes | P2 estimate reflects that most minor underdoses are not symptomatic or provide ample opportunity for users to correct before significant impact
Dose | Delay in therapy (device not used) | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact | 1 | Yes | Assumption: dose is replaced within reasonable time
Drug | Wrong drug delivered (quantity or characteristics associated with low toxicity) | 4 | Less than full therapeutic benefit (likely no clinical safety concerns) | 2 | Yes | P2 estimate reflects that most situations will provide ample opportunity for users to correct before significant impact and variations in victim sensitivity/resistance and wrong drug delivered.
Drug | Wrong drug delivered (quantity or characteristics associated with high toxicity) | 4 | Drug dependent SERIOUS adverse event (e.g.., anaphylaxis), possibly requiring major medical intervention | 4 | Yes | P2 estimate reflects the wide expected range of victim sensitivity/ resistance and variation in wrong drug.

## PHA
Hazard | Hazardous Situation | P2 | Harm | Severity | Medication | Regulatory Review | Comments | Taltz | Omvoh
No Hazard | N/A - other situation with no associated hazard | N/A | N/A – no clinical impact | N/A | N/A | No | None
User | User dissatisfaction | 1 | None | 1 | Taltz | No | None
User | User dissatisfaction | 1 | None | 1 | Omvoh | No | None
Dose Logging | Dose that was not taken shows up in the logbook | 4 | Less than full therapeutic benefit (Underdose) | 2 | Taltz | Yes | P2 estimate reflects that most minor underdoses are not symptomatic or provide ample opportunity for users to correct before significant impact Drug harm from: Ixe 1mL ePFS System RMF (118)
Dose Logging | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz | Yes | P2 estimate reflects that most minor overdoses are not symptomatic or provide ample opportunity for users to correct before significant impact Drug harm from: Ixe 1mL ePFS System RMF (118)
Dose Logging | Incorrect scheduled dosing day is identified | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz | Yes | P2 estimate reflects that most minor overdoses are not symptomatic or provide ample opportunity for users to correct before significant impact Drug harm from: Ixe 1mL ePFS System RMF (118)
Dose Logging | Dose that was not taken shows up in the logbook | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication) | 3 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
Dose Logging | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
Dose Logging | Incorrect scheduled dosing day is identified | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
Dose Reminder | No medication reminder is sent | 4 | Less than full therapeutic benefit (Underdose) | 2 | Taltz | Yes | P2 estimate reflects that most minor underdoses are not symptomatic or provide ample opportunity for users to correct before significant impact Drug harm from: Ixe 1mL ePFS System RMF (118)
Dose Reminder | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz | Yes | P2 estimate reflects that most minor overdoses are not symptomatic or provide ample opportunity for users to correct before significant impact Drug harm from: Ixe 1mL ePFS System RMF (118)
Dose Reminder | No medication reminder is sent | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication) | 3 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
Dose Reminder | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
App Configuration | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz | Yes | P2 estimate reflects that most minor overdoses are not symptomatic or provide ample opportunity for users to correct before significant impact Drug harm from: Ixe 1mL ePFS System RMF (118)
App Configuration | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 1 | Taltz | Yes | Assumption: dose is replaced within reasonable time Drug harm from: Ixe 1mL ePFS System RMF (118)
App Configuration | Connected Auto Injector functionality is enabled | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz | Yes | P2 estimate reflects that most minor overdoses are not symptomatic or provide ample opportunity for users to correct before significant impact Drug harm from: Ixe 1mL ePFS System RMF (118)
App Configuration | Application UI does not load properly | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 1 | Taltz | Yes | Assumption: dose is replaced within reasonable time Drug harm from: Ixe 1mL ePFS System RMF (118)
App Configuration | Application UI displays incorrect content or feature | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication) | 3 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
App Configuration | Registered user is unable to use the application | 5 | Delay in Therapy (device not used) | 1 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
App Configuration | Connected Auto Injector functionality is enabled | 4 | Overdose up to 350% (absolute dose < 900 mg) | 2 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
App Configuration | Application UI does not load properly | 5 | Delay in Therapy (device not used) | 1 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
Symptom tracker | Inaccurate information is displayed in the symptom tracker | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz | Yes | P2 estimate reflects that most minor overdoses are not symptomatic or provide ample opportunity for users to correct before significant impact Drug harm from: Ixe 1mL ePFS System RMF (118)
Symptom tracker | Inaccurate information is displayed in the symptom tracker | 4 | Overdose up to 350% (absolute dose < 900 mg) | 2 | Omvoh | Yes | Drug harm from: Tella Miri System RMF (185)
Symptom tracker | Inaccurate information is displayed in the symptom tracker | 4 | Moderate: State of injury that is not permenant, but may require minor medical intervention (Overdose > 40mg (use related)) | 2 | Olumiant | Yes | Drug Harm from: Stature - Baricitinib Container Closure System PHA Revision 2
Note on A23: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: @Jill E Vaughn and @Heydy Andolz- The team was wanting your thoughts as to whether or not a problem with the Symptom Tracker can actually lead to a patient taking more drug. Can you please confirm? Should we keep these as is or remove them? I can set up some time to discuss if needed.

## RiskLookUp
P2
| P harm |  | Remote | Low | Moderate | High | Very high
|  |  | 1 | 2 | 3 | 4 | 5 | [5 empty] | SEV
P1 | Very high | 5 | 1 | 2 | 3 | 4 | 5 |  |  |  | negligible | minor | moderate | major | severe
| High | 4 | 1 | 1 | 2 | 3 | 4 |  |  |  | 1 | 2 | 3 | 4 | 5
| Moderate | 3 | 1 | 1 | 2 | 2 | 3 | Ph | Very high | 5 | Low 1-5 | Medium 2-5 | High 3-5 | High 4-5 | High 5-5
| Low | 2 | 1 | 1 | 1 | 1 | 2 |  | High | 4 | Low 1-4 | Medium 2-4 | Medium 3-4 | High 4-4 | High 5-4
| Remote | [7 empty] | Moderate | 3 | Low 1-3 | Low 2-3 | Medium 3-3 | Medium 4-3 | High 5-3 | bar: 5 to 5
[9 empty] | Low | 2 | Low 1-2 | Low 2-2 | Low 3-2 | Medium 4-2 | Medium 5-2
[9 empty] | Remote | 1 | Low 1-1 | Low 2-1 | Low 3-1 | Low 4-1 | Medium 5-1
Table 4. Probability Ranking
| Category | Qualitative Description | FMEA Ranking | P1 | P2 | P-Harm
|  |  |  | (events/patient year of therapy) |  | (events/patient year of therapy)
| Frequent | Common, repeated, or occurring often | 5 | > 1/100 | > 1/100 | > 1/100
| Probable | Likely to occur | 4 | ≤ 1/100 | ≤ 1/100 | ≤ 1/100
| Occasional | Occurring at irregular or infrequent intervals | 3 | ≤ 1/1,000 | ≤ 1/1,000 | ≤ 1/1,000
| Remote | Unlikely, but can reasonably be expected to occur | 2 | ≤ 1/100,000 | ≤ 1/100,000 | ≤ 1/100,000
| Improbable | Unlikely to occur; not expected | 1 | ≤ 1/10,000,000 | ≤ 1/10,000,000 | ≤ 1/10,000,000

## Heat Map
SEV
|  |  | negligible | minor | moderate | major | severe
|  |  | 1 | 2 | 3 | 4 | 5
Ph | Very high | 5 | 17 | 0 | 0 | 0 | 0
| High | 4 | 0 | 0 | 0 | 0 | 0
| Moderate | 3 | 2 | 1 | 0 | 0 | 0
| Low | 2 | 0 | 9 | 0 | 0 | 0
| Remote | 1 | 0 | 10 | 0 | 0 | 0
Risk Summary
|  |  |  | High | 0
|  |  |  | Medium | 0
|  |  |  | Low | 39
|  |  |  | Total | 39

## List of System Effects
System Effects
N/A - No System Effect
App logs injection event from a device that does not belong to the user
App presents incorrect logbook data
App does not log successful injection event
App logs unsuccessful injection event as injection event
App logs partial injection event as injection event
App presents multiple device logbook entries including the users device
A registered user is unable to use the mobile application.
An unregistered user is unable to use the mobile application.
Unintended user is able to access dose tracking functionality
Application displays an incorrect scheduled dosing day
Dose reminder is sent on incorrect day
Dose reminder is sent at incorrect time
Incorrect configuration is used causing incorrect features or content to be displayed
Dose reminder is not sent
Incorrect reminders are sent
User's tracked symptom photos are stored in the device camera roll
User-scheduled appointment reminder is not sent

## Revision History
Revisions
Updated text throughout sFMEA tab to remove references to specific drugs (Taltz) except for references to the Connected Taltz Delivery System
Updated controls RCM-021, RCM-022, RCM-023 to remove references to Taltz since these controls apply to Omvoh as well
Updated "Function Status" for Automatic Injection Tracking feature status in R11
Updated RCM-024 to be more generic to meet multiple use cases
Added control RCM-001 to symptom tracking
Updated several system effects to standardize language across risk items
Updated RCM-031, RCM-033, RCM-034 to refer to medication administrations generally instead of injections, so this control can be referenced for multiple drugs
Added Omvoh SFMEA
Added new features Accident Tracking and Infusion Tracking

## Taltz PHA Diff
Hazard (R10) | Hazardous Situation | P2 | Harm | Severity | Medication | Hazard (R11) | Hazardous Situation | P2 | Harm | Severity | Medication
User Experience | User dissatisfaction | 5 | User dissatisfied | 1 | Taltz | Dose Logging | Dose that was not taken shows up in the logbook | 4 | Less than full therapeutic benefit (Underdose) | 2 | Taltz
No Hazard | N/A - other situation with no associated hazard | N/A | N/A – no clinical impact | N/A | N/A |  | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
Dose | Repeated missed dose or partial dose for up to 3 months | 4 | Minor: Less than full therapeutic benefit, possible requiring medical intervention | 3 | Taltz |  | Incorrect scheduled dosing day is identified | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
Dose | Delivery of drug with less than full potency repeated for up to 3 months (equivalent to underdose) | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention | 3 | Taltz |  | Dose that was not taken shows up in the logbook | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication) | 3 | Omvoh
Dose | Underdose | 4 | Minor: Less than full therapeutic benefit | 2 | Taltz |  | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh
Dose | No dose delivered | 4 | Minor: Less than full therapeutic benefit | 2 | Taltz |  | Incorrect scheduled dosing day is identified | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh
Dose | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | Taltz | Dose Reminder | No medication reminder is sent | 4 | Less than full therapeutic benefit (Underdose) | 2 | Taltz
Dose | Delivery of drug with less than full potency (equivalent to underdose) | 4 | Minor: Less than full therapeutic benefit | 2 | Taltz |  | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
Dose | Delay in therapy (device not used) | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact | 1 | Taltz |  | No medication reminder is sent | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication) | 3 | Omvoh
Drug | Wrong drug delivered (quantity or characteristics associated with low toxicity) | 4 | Less than full therapeutic benefit (likely no clinical safety concerns) | 2 | Taltz |  | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh
Drug | Wrong drug delivered (quantity or characteristics associated with high toxicity) | 4 | Drug dependent SERIOUS adverse event (e.g.., anaphylaxis), possibly requiring major medical intervention | 4 | Taltz | App Configuration | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
[7 empty] | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 1 | Taltz
[7 empty] | Connected Auto Injector functionality is enabled | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
[7 empty] | Application UI does not load properly | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 1 | Taltz
[7 empty] | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Omvoh
[7 empty] | Registered user is unable to use the application | 5 | NEGLIGIBLE State of dissatisfaction with no perceptible clinical impact (Delay in Therapy (device not used)) | 1 | Omvoh
[7 empty] | Connected Auto Injector functionality is enabled | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Omvoh
[7 empty] | Application UI does not load properly | 5 | NEGLIGIBLE State of dissatisfaction with no perceptible clinical impact (Delay in Therapy (device not used)) | 1 | Omvoh
[6 empty] | Symptom tracker | Inaccurate information is displayed in the symptom tracker | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
[7 empty] | Inaccurate information is displayed in the symptom tracker | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Omvoh
[7 empty] | Inaccurate information is displayed in the symptom tracker | 4 | Minor: Temporary and non-debilitating state of pain/discomfort (Overdose < 40mg (use related)) | 2 | Olumiant
[6 empty] | No Hazard | N/A - other situation with no associated hazard | N/A | N/A | N/A – no clinical impact | Taltz
Note on H21: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: @Jill E Vaughn and @Heydy Andolz- The team was wanting your thoughts as to whether or not a problem with the Symptom Tracker can actually lead to a patient taking more drug. Can you please confirm? Should we keep these as is or remove them? I can set up some time to discuss if needed.
NOTES
1. The R11 PHA does not have a Harm = User dissatisfied. There are 17 instances of this in the sFMEA.
2. Minor: Less than full therapeutic benifit - In the R10 PHA there are two'Minor: Less than full therapeutic benifit' harms. One is associated with an Underdose Haz Sit and the other is associated with the No Dose Delivered Haz Sit. There are 3 instances of the Underdose and no instances of the No dose delivered.
3. Any R10 PHA harm not linked to R11 PHA were not referenced in the sFMEA
4. Linked to Harms vs Haz Sits.
5. There are three R11 harms that match the R10 harms if you combine the R10 harm language with the R10 Haz Sit language (ref Row 7, 9, and 11).
R10 Harms for Taltz

## Omvoh PHA Diff
Hazard | Hazardous Situation | P2 | Harm | Severity | Medication | Hazard (R11) | Hazardous Situation | P2 | Harm | Severity | Medication
User Experience | User dissatisfaction | 5 | User dissatisfied | 1 | Omvoh | Dose Logging | Dose that was not taken shows up in the logbook | 4 | Less than full therapeutic benefit (Underdose) | 2 | Taltz
No Hazard | N/A - other situation with no associated hazard | N/A | N/A – no clinical impact | N/A | N/A |  | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
Dose | Repeated missed dose or partial dose for up to 3 months | 4 | Minor: Less than full therapeutic benefit, possible requiring medical intervention | 3 | Omvoh |  | Incorrect scheduled dosing day is identified | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
Dose | Delivery of drug with less than full potency repeated for up to 3 months (equivalent to underdose) | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention | 3 | Omvoh |  | Dose that was not taken shows up in the logbook | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication) | 3 | Omvoh | Harm different than Taltz
Dose | Underdose | 4 | Minor: Less than full therapeutic benefit | 2 | Omvoh |  | Dose that was taken does not show up in the logbook | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh | Harm different than Taltz The absolute dose < 900 version is not in the R10 PHA
Dose | No dose delivered | 4 | Minor: Less than full therapeutic benefit | 2 | Omvoh |  | Incorrect scheduled dosing day is identified | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh | Harm different than Taltz The absolute dose < 900 version is not in the R10 PHA
Dose | Overdose up to 200% (80-160 mg absolute error) | 4 | Minor: pain/discomfort | 2 | Omvoh | Dose Reminder | No medication reminder is sent | 4 | Less than full therapeutic benefit (Underdose) | 2 | Taltz
Dose | Delivery of drug with less than full potency (equivalent to underdose) | 4 | Minor: Less than full therapeutic benefit | 2 | Omvoh |  | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
Dose | Delay in therapy (device not used) | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact | 1 | Omvoh |  | No medication reminder is sent | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication) | 3 | Omvoh | Harm different than Taltz
Drug | Wrong drug delivered (quantity or characteristics associated with low toxicity) | 4 | Less than full therapeutic benefit (likely no clinical safety concerns) | 2 | Omvoh |  | Dose reminder is sent at the wrong time | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 350% (absolute dose < 900 mg)) | 2 | Omvoh | Harm different than Taltz The absolute dose < 900 version is not in the R10 PHA
Drug | Wrong drug delivered (quantity or characteristics associated with high toxicity) | 4 | Drug dependent SERIOUS adverse event (e.g.., anaphylaxis), possibly requiring major medical intervention | 4 | Omvoh | App Configuration | Application UI displays incorrect content or feature | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
[7 empty] | Registered user is unable to use the application | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 1 | Taltz
[7 empty] | Connected Auto Injector functionality is enabled | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
[7 empty] | Application UI does not load properly | 5 | Negligible: state of dissatisfaction with no perceptible clinical impact (Delay in therapy (device not used)) | 1 | Taltz
[7 empty] | Application UI displays incorrect content or feature | 4 | MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication) | 3 | Omvoh | CHANGE FROM: MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) TO: MODERATE Less than full therapeutic benefit, possibly requiring medical intervention (Underdose 5%-100% - UC or Crohn's indication)
[7 empty] | Registered user is unable to use the application | 5 | NEGLIGIBLE State of dissatisfaction with no perceptible clinical impact (Delay in Therapy (device not used)) | 1 | Omvoh
[7 empty] | Connected Auto Injector functionality is enabled | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Omvoh
[7 empty] | Application UI does not load properly | 5 | NEGLIGIBLE State of dissatisfaction with no perceptible clinical impact (Delay in Therapy (device not used)) | 1 | Omvoh
[6 empty] | Symptom tracker | Inaccurate information is displayed in the symptom tracker | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Taltz
[7 empty] | Inaccurate information is displayed in the symptom tracker | 4 | MINOR pain/ discomfort (no clinical safety concerns) (Overdose up to 200% (80-160 mg absolute error)) | 2 | Omvoh
[7 empty] | Inaccurate information is displayed in the symptom tracker | 4 | Minor: Temporary and non-debilitating state of pain/discomfort (Overdose < 40mg (use related)) | 2 | Olumiant | Harm different than Taltz and Omvoh
[6 empty] | No Hazard | N/A - other situation with no associated hazard | N/A | N/A | N/A – no clinical impact | Taltz
R10 Harms for Omvoh
NOTES
1. The R11 PHA does not have a Harm = User dissatisfied. There are 17 instances of this in the sFMEA.
2. Minor: Less than full therapeutic benifit - In the R10 PHA there are two'Minor: Less than full therapeutic benifit' harms. One is associated with an Underdose Haz Sit and the other is associated with the No Dose Delivered Haz Sit. There are 3 instances of the Underdose and no instances of the No dose delivered.
3. Any R10 PHA harm not linked to R11 PHA were not referenced in the sFMEA
4. Linked to Harms vs Haz Sits.
5. There are two R11 harms that match the R10 harms if you combine the R10 harm language with the R10 Haz Sit language (ref Row 9 and 11).
6. For the R11 PHA Omvoh has two 'MINOR pain/discomfort' harms. One has an absolute error of 80-160 mg and the other has an absolute error of < 900 mg. The R10 PHA has only the 80-160 mg version. Is it possible that some of the Omvoh items will need to change to the < 900 mg absolute error?
Note on H20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: @Jill E Vaughn and @Heydy Andolz- The team was wanting your thoughts as to whether or not a problem with the Symptom Tracker can actually lead to a patient taking more drug. Can you please confirm? Should we keep these as is or remove them? I can set up some time to discuss if needed.